# Prohibited ingredients in vaping products by Gas Chromatography-Mass Spectrometry

June 2025

Copyright

© Commonwealth of Australia 2025
This work is copyright. You may reproduce the whole or part of this work in unaltered form for your own personal use or, if you are part of an organisation, for internal use within your organisation, but only if you or your organisation do not use the reproduction for any commercial purpose and retain this copyright notice and all disclaimer notices as part of that reproduction. Apart from rights to use as permitted by the Copyright Act 1968 or allowed by this copyright notice, all other rights are reserved and you are not allowed to reproduce the whole or any part of this work in any way (electronic or otherwise) without first being given specific written permission from the Commonwealth to do so. Requests and inquiries concerning reproduction and rights are to be sent to the TGA Copyright Officer, Therapeutic Goods Administration, PO Box 100, Woden ACT 2606 or emailed to <tga.copyright@tga.gov.au>.

Contents

Instrument set up	4

Chromatographic system	4

MRM Mode: Mass Spectrometer System	4

Limits of detection	5

Solutions	5

Prohibited ingredients stock solution	5

Working Standard	5

Quality control sample	6

Limit standard	6

Concentrated standard	6

Sample solution	6

Injection sequence	6

System suitability criteria for the Limit solution	7

Identification criteria	7

Prohibited Ingredients:	7

Prohibited ingredients in vaping products by GCMS

This method describes the identification of the following prohibited ingredients in vaping products by Gas Chromatography-Mass Spectrometry (GCMS): Diacetyl; 2,3-Pentanedione; Acetoin; Ethylene glycol; Benzaldehyde; Diethylene glycol and trans-Cinnamaldehyde.

## Instrument set up

### Chromatographic system

Table 1: Chromatographic system set up

| Instrument: | GCMS |  |  |
| --- | --- | --- | --- |
| Column: | SH-624, 0.32 mm ID, 1.8 µm DF and 30 m L. |  |  |
| Oven Temperature Program: | Rate | Temperature | Hold Time |
|  | n/a | 40 ºC | 3.00 min |
|  | 10 ºC/min | 210 ºC | 0.00 min |
|  | 25 ºC/min | 300 ºC | 1.40 min |
| Injection Temperature | 250 ºC |
| Injection Mode | Split |
| Split Ratio | 20.0 |
| Column Flow | 2.00 mL/min |

### MRM Mode: Mass Spectrometer System

Table 2: Ion source and interface temperatures

| Ion Source Temperature | 200 ºC |
| --- | --- |
| Interface Temperature | 250 ºC |

Table 3: MRM event table for method: Quantifier transitions

| Peak | Start time | End time | Precursor → Product | CE |
| --- | --- | --- | --- | --- |
| Diacetyl | 3.70 | 5.50 | 86→43 | 6 |
| 2,3-pentanedione | 5.50 | 6.84 | 100→57 | 6 |
| Acetoin | 6.84 | 7.52 | 45→27 | 12 |
| Ethylene glycol | 7.52 | 10.00 | 62→33 | 3 |
| Benzaldehyde | 10.00 | 12.59 | 106→105 | 6 |
| Diethylene glycol | 12.59 | 15.80 | 75→45 | 6 |
| Trans-Cinnamaldehyde | 17.38 | 18.18 | 131→77 | 27 |

Table 4: MRM event table for method: Qualifier transitions

| Peak | Precursor → Product | CE | Ion ratio | Precursor → Product | CE | Ion ratio1 |
| --- | --- | --- | --- | --- | --- | --- |
| Diacetyl | 86→86 | 0 | 3.0 ± 1.5 | n/a | n/a | n/a |
| 2,3-pentanedione | 100→43 | 12 | 9.4 ± 2.8 | n/a | n/a | n/a |
| Acetoin | 45→29 | 9 | 106 ± 32 | n/a | n/a | n/a |
| Ethylene glycol | 62→31 | 6 | 55 ± 16 | n/a | n/a | n/a |
| Benzaldehyde | 106→77 | 16.5 | 50 ± 15 | 106→78 | 15 | 4.9 ± 2.4 |
| Diethylene glycol | 45→27 | 12 | 18 ± 6 | 45→29 | 9 | 15 ± 4 |
| Trans-Cinnamaldehyde | 131→103 | 9 | 45 ± 14 | 131→131 | 0 | 18 ± 5 |

### Limits of detection

Table 5: Analyte Limits of detection

| Analyte | Limit of Detection (μg/mL) |
| --- | --- |
| Diacetyl | 0.3 |
| 2,3-pentanedione | 0.3 |
| Acetoin | 0.2 |
| Ethylene glycol | 0.8 |
| Benzaldehyde | 0.2 |
| Diethylene glycol | 0.8 |
| Trans-Cinnamaldehyde | 0.2 |

## Solutions

### Prohibited ingredients stock solution

Prepare individual solutions in acetonitrile of the following analytes with the concentrations described in the Table 6. These solutions can be stored at 4 ºC in an amber flask for up to 40 days.

Table 6: Analyte stock solution concentrations in acetonitrile

| Analyte | Concentration (μg/mL) |
| --- | --- |
| Diacetyl | 1000 |
| 2,3-Pentanedione | 1000 |
| Acetoin | 1000 |
| Benzaldehyde | 1000 |
| trans-Cinnamaldehyde | 1000 |
| Ethylene glycol | 2500 |
| Diethylene glycol | 2500 |

### Working Standard

Prepare a mixed standard using the stock solutions and methanol as the diluent, adjusting each individual aliquot to obtain the concentrations in Table 7. These solutions should be made fresh prior to testing.

Table 7: Analyte stock solution concentrations in acetonitrile

| Analyte | Concentration (μg/mL) |
| --- | --- |
| Diacetyl | 10 |
| 2,3-Pentanedione | 10 |
| Acetoin | 10 |
| Benzaldehyde | 10 |
| trans-Cinnamaldehyde | 10 |
| Ethylene glycol | 25 |
| Diethylene glycol | 25 |

### Quality control sample

Prepare a mixture of glycerol and propylene glycol by measuring 7 mL of glycerol and 3 mL of propylene glycol and mixing the two until a homogeneous mixture is obtained (QC sample).

Using a suitable positive displacement piston operated volumetric apparatus (POVA), transfer 80 µL of the resulting mixture to an HPLC vial and then transfer 1520 µL of methanol (QC sample blank).

### Limit standard

Using a suitable positive displacement POVA, transfer 80 µL of the QC sample to an HPLC vial and then transfer 80 µL of the Working Standard followed by addition of 1440 µL of methanol (QC sample – limit test solution).

### Concentrated standard

Using a suitable positive displacement POVA, transfer 80 µL of the QC sample to an HPLC vial and then transfer 1520 µL of the Working Standard (Second identification standard).

### Sample solution

Prepare a fresh 1 in 20 dilution of the sample in methanol by transferring 80 µL of the sample to an HPLC vial followed by 1520 µL of methanol using a suitable positive displacement POVA.

For samples that contain a large amount of acetoin which is not identified due to the ion ratios being saturated, a further dilution can be prepared to improve the ion ratio match.

## Injection sequence

Inject in this order the following solutions:

Methanol	2 injections

QC Sample blank	2 injections

Limit Standard	2 injections

Concentrated standard	2 injections

Samples (up to 6 samples)	1 injection

Methanol	1 injection

QC Sample blank	1 injection

Limit Standard	1 injection

Concentrated standard	1 injection

Samples (up to 6 samples)	1 injection

Methanol	1 injection

QC Sample blank	1 injection

Limit Standard	1 injection

Concentrated standard	1 injection

End

### System suitability criteria for the Limit solution

The %RSD of the area of each analyte for the Limit solution (injected at least 3 times) should be NMT 20%.

## Identification criteria

#### Prohibited Ingredients:

Disregard any peaks with areas below the areas observed in the Limit Standard.

Disregard any peaks with retention time difference (absolute) more than 0.10 min compared with the same analyte in the Limit Standard.

Disregard any peaks where the ion ratio test fails, taking into consideration the ion ratios set up either by the Limit Standard or the Concentrated standard for identification.

A prohibited ingredient is present when the area of the peak is greater than the corresponding peak in the limit standard, and the retention time difference and ion ratio requirements are met.

The ion ratios in this table are indicative only, these should be updated using the Limit Standard or the Concentrated standard. Specifications: Ion ratio ≤ 5 (± 50% relative); Ion ratio > 5 (± 30% relative).