## TGA
AUSTRALIAN REGIONAL nIVD ToC MATRIX - LEGEND R = Required CR = Conditionally Required O = Optional OR = Optional but Recommended NR = Not Required - = No Content (requirement should be based on the sub heading classification)
|  |  | Note: for Content required under each Heading see the IMDRF Table of Contents document | Conformity Assessment Application - Procedures under Schedule 3, Part 1 or Part 4 |  | Regional content - Australia - Therapeutic Goods Administration
Row ID | Scope | Level | Heading | Classification | Condition
CHAPTER 1 – REGIONAL ADMINISTRATIVE
CH1.01 | IMDRF, RF | 1 | Cover Letter | R |  | In addition to IMDRF content, the cover letter of the application needs to be prepared on company letterhead and to also include: a) Submission ID that is generated electronically when completing the application form in TBS b) Contact details of the person authorised to liaise with TGA during the evaluation process c) Whether an abridged assessment is being sought, and the basis for making this request d) Signature of the authorised person for the company
CH1.02 | IMDRF | 1 | Submission ToC | R |  | Same as IMDRF content
CH1.03 | IMDRF | 1 | List of Terms/Acronyms | R |  | Same as IMDRF content
CH1.04 | Regional | 1 | Application Form/Administrative Information | R |  | The TBS Application form in pdf to include administrative data of the applicant (including Client ID, contact person and details), application scope (including applicable conformity assessment procedure and type of application (new, change, recertification, current certification details, manufacturer details (including Client ID, contact details and facility scope), and device details including GMDN code, term and classification. Refer to www.tga.gov.au for the most up to date information.
CH1.05 | IMDRF, RF (ANVISA, EU, TGA, USFDA) | 1 | Listing of Device(s) | R |  | In addition to IMDRF content, for all classes of devices the applicant needs to include: a) The Global Medical Device Nomenclature (GMDN) Code and Term b) The classification and the applicable classification rule c) Device description and intended purpose For class III and AIMDs this table should also identify the following: a) Unique Product Identifiers (see the Therapeutic Goods (Medical Devices) Regulations 2002) b) Variants, if applicable (as defined in the Therapeutic Goods (Medical Devices) Regulations 2002)
CH1.06 | Regional (HC, Anvisa, EU, TGA) | 1 | Quality Management System, Full Quality System or Product Certification Certificate | CR | If applicable | Copies of any current TGA or other regulatory authority certification referenced within the submission or required for the submission type to be considered for reduced assessment. Refer to TGA Guidance "Reduction of assessment fees for Medical Devices".
CH1.07 | Regional (Anvisa, CFDA) | 1 | Free Sale Certificate | NR
CH1.08 | Regional (USFDA, Anvisa, EU) | 1 | User Fees | NR
CH1.09 | IMDRF, RF | 1 | Pre-Submission Correspondence and Previous Regulator Interactions | R |  | Same as IMDRF content
CH1.10 | Regional (USFDA, TGA) | 1 | Acceptance for Review Checklist | NR | NOTE: This was previously shown as required for TGA but ToC submission replaces the Supporting Data Checklist
CH1.11 | Regional (ANVISA, HC, EU, TGA, USFDA) | 1 | Statements/Certifications/Declarations of Conformity | -
CH1.11.1 | Regional (USFDA) | 2 | Performance and Voluntary Standard | NR
CH1.11.2 | Regional (USFDA) | 2 | Environmental Assessment | NR
CH1.11.3 | Regional (USFDA) | 2 | Clinical Trial Certifications | NR
CH1.11.4 | Regional (USFDA) | 2 | Indications for Use Statement with Rx and OTC designation Enclosure | NR
CH1.11.5 | Regional (ANVISA, CFDA, HC, TGA, USFDA) | 2 | Truthful and Accurate Statement | R | Manufacturer's statutory declaration | TGA Conformity Assessment - Manufacturer's statutory declaration a) A statutory declaration is a written statement allowing a person to declare something to be true. The declaration is signed in the presence of a witness. Giving false or misleading information as part of a statutory declaration is a criminal offence under the Criminal Code. Refer: https://www.tga.gov.au/form/manufacturer-statutory-declarations The following statements of undertaking by the manufacturer are required by conformity assessment procedure(s) set in the Therapeutic Goods (Medical Devices) Regulations 2002: • An undertaking(in writing) by the manufacturer to continue to comply with the requirements of the quality management system after assessment (Regulations - Schedule 3, Part 1.3(2)(e), Part 4.3(2)(e)). • An undertaking(in writing) by the manufacturer to ensure that the quality management system is at all times adequate and efficacious (Regulations - Schedule 3, Part 1.3(2)(f), Part 4.3(2)(f)). • An undertaking(in writing) by the manufacturer to notify the Secretary or the Australian sponsor, of any information of the kind mentioned in subparagraphs 1.4(3)(c)(i) or (ii),4.4(3)(c)(i) or (ii), that the manufacturer becomes aware of in relation to the kind of medical device (Regulations - Schedule 3, Part 1.3(2)(g), Part 4.3(2)(i)).
CH1.11.6 | Regional (USFDA) | 2 | USFDA Class III Summary and Certification | NR
CH1.11.7 | IMDRF (EU, JP, TGA) | 2 | Declaration of Conformity | R | Draft DoC | In addition to IMDRF content, the wording of the Declaration of Conformity will depend on the conformity assessment procedure chosen by the manufacturer. Templates for each of the six possible types of Declarations of Conformity under Schedule 3 of the Therapeutic Goods (Medical Devices) Regulations 2002 are available at https://www.tga.gov.au/form/declaration-conformity-templates-medical-devices.
CH1.12 | IMDRF | 1 | Letters of Reference for Master Files | CR | When a Master File is referenced. | For devices containing new chemical(medicinal substance) entity(s): if a DMF/PMF and/or EDQM Certificate of Suitability is submitted to TGA to demonstrate quality compliance then it is required that the DMF/PMF reference number along with the Letter of authorisation to access the master file of the medicinal substance be provided (i.e. a letter from the medicinal substance manufacturer to allow TGA access to the DMF). For devices containing non-new chemical entities manufactured by a different manufacturer that is not on the ARTG (Australian Register of Therapeutic Goods): If a DMF/PMF and /or EDQM Certificate of Suitability is submitted to TGA to demonstrate quality compliance then it is required that the DMF/PMF reference number along with the Letter of access of authorisation to the master file of the medicinal substance be provided (i.e. a letter from the medicinal substance manufacturer to allow TGA access to the DMF). The template to provide letter of access to DMF/PMF and /or EDQM Certificate of Suitability can be obtained from Common Technical Document 1.6 (https://www.tga.gov.au/book/16-master-files-and-certificates-suitability). For additional information refer to Guidance 11 of ARGPM. (https://www.tga.gov.au/publication/australian-regulatory-guidelines-prescription-medicines-argpm ).
CH1.13 | Regional (ANVISA, CFDA) | 1 | Letter of Authorization | NR
CH1.14 | IMDRF | 1 | Other Regional Administrative Information | CR | If applicable | Same as IMDRF content
CHAPTER 2 – SUBMISSION CONTEXT
CH2.1 | IMDRF | 1 | Chapter ToC | R |  | Same as IMDRF content
CH2.2 | IMDRF, RF | 1 | General Summary of Submission | R | If application is for recertification or change to CA certificate, relevant TGA certicate numbers must be detailed | Same as IMDRF content
CH2.3 | Regional (USFDA) | 1 | Summary and Certifications for Premarket Submissions | NR
CH2.4 | IMDRF | 1 | Device Description | -
CH2.4.1 | IMDRF, RF | 2 | Comprehensive Device Description & Principle of Operation | R |  | In addition to IMDRF content, the device description should also include the following information: Does the device incorporate a measuring function? Does the device incorporate medicinal substance, human blood or plasma that has an action ancillary to the device? Does the device incorporate materials of animal/microbial/recombinant origin?
CH2.4.2 | IMDRF (ANVISA, CFDA, EU, HC, TGA, USFDA) | 2 | Description of Device Packaging | R |  | Same as IMDRF content
CH2.4.3 | IMDRF | 2 | History of Development | CR | Required for devices undergoing design or type examination | Same as IMDRF content
CH2.4.4 | IMDRF, RF | 2 | Reference and Comparison to Similar and/or Previous Generations of the Device | CR | Required when relevent and in support of the application (i.e. Testing leveraged from previous generation of device etc...) | In addition to IMDRF content, in regards to clinical data requirements: a) Include a comparison of indications for use and the technology (including features materials and principles of operation) between the predicate device(s) and subject device(s). b) Include an analysis of why any differences between the subject device(s) and the predicate device(s) do not render the subject device(s) Not Substantially Equivalent, affect safety or effectiveness or raise different questions of safety and effectiveness.
CH2.4.5 | Regional (USFDA) | 2 | Substantial Equivalence Discussion | NR
CH2.5 | IMDRF, RF | 1 | Indications for Use and/or Intended Use and Contraindications | -
CH2.5.1 | IMDRF, RF | 2 | Intended Use, Intended Purpose, Intended User, Indications for Use | R |  | Same as IMDRF content
CH2.5.2 | IMDRF, RF | 2 | Intended Environment/Setting for use | R |  | The information in the IMDRF content, should all be included in the IFU or product labelling.
CH2.5.3 | Regional (USFDA) | 2 | Pediatric Use | NR
CH2.5.4 | IMDRF, RF | 2 | Contraindications for Use | CR | If applicable | In addition to IMDRF content, this information should be included in the IFU or product labelling.
CH2.6 | IMDRF | 1 | Global Market History | -
CH2.6.1 | IMDRF | 2 | Global Market History | R |  | In addition to IMDRF content, any notifications to foreign regulators of substantial change to the device.
CH2.6.2 | IMDRF, RF | 2 | Global Incident Reports and Recalls | R |  | Same as IMDRF content
CH2.6.3 | IMDRF, RF (HC, JP, EU, TGA) | 2 | Sales, Incident and Recall Rates | R |  | Same as IMDRF content
CH2.6.4 | Regional (TGA) | 2 | Evaluation/Inspection Reports | CR | When information is relevant in support of the application, e.g. Notified Body reports | Copies of Evaluation/Inspection Reports from other parties (e.g. Notified Body audit reports). • The audit reports provided should be from the most recent audit and if the most recent audit is not a full initial or recertification audit, then the report from the most recent full or recertification audit should also be provided. • The technical assessment reports should be relevant to the subject device in the submission. For example: Application for change to a device should provide technical assessment report relevant to the change referred to in the submission.
CH2.7 | IMDRF | 1 | Other Submission Context Information | CR | When information is requested by the regulator (through guidance documents or other communication) but does not belong in any of the other submission context headings. | Same as IMDRF content
CHAPTER 3 – NON-CLINICAL EVIDENCE
CH3.1 | IMDRF | 1 | Chapter ToC | R
CH3.2 | IMDRF | 1 | Risk Management | R |  | Same as IMDRF content
CH3.3 | IMDRF (ANVISA, EU, JP, TGA) | 1 | Essential Principles (EP) Checklist | R |  | Same as IMDRF content
CH3.4 | IMDRF, RF | 1 | Standards | -
CH3.4.1 | IMDRF, RF | 2 | List of Standards | CR | Required unless this information is presented in the Essential Principles Checklist. | This list should include any medical device standard or conformity assessment standard that has been applied to the device; and, if no medical device standard or conformity assessment standard, or part only of such a standard, has been applied to the device — the solutions adopted to ensure that each device complies with the applicable provisions of the essential principles. The information in this section may be presented in the Essential Principle Checklist and, if so, needs only to be presented once in the application.
CH3.4.2 | Regional (HC, Anvisa, USFDA) | 2 | Declaration and/or Certification of Conformity | NR
CH3.5 | IMDRF | 1 | Non-clinical Studies | -
CH3.5.01 | IMDRF | 2 | Physical and Mechanical Characterization | CR | If applicable | In addition to IMDRF content, where applicable, testing certificate and accreditation status of laboratories used for physical and mechanical parameters testing in accordance with relevant applicable standards needs to be provided.
CH3.5.01.1 | IMDRF | 3 | [Study description, study identifier, date of initiation] | CR | When any testing/study is conducted in support of the submission under the parent heading subject, this is required for each study/test. See information on Numbering of Headings in IMDRF ToC document
CH3.5.01.1.1 | IMDRF | 4 | Summary | CR | If applicable | Same as IMDRF content
CH3.5.01.1.2 | IMDRF | 4 | Full Report | CR | If applicable | Same as IMDRF content
CH3.5.01.1.3 | Regional (USFDA) | 4 | Statistical Data | NR
CH3.5.02 | IMDRF | 2 | Chemical /Material Characterization | CR | If applicable | In addition to IMDRF content, this should also include a list of all chemicals/materials used in the device along with the evidence of quality of all of the raw materials (e.g. Certificate of Analysis). If the device incorporates a new chemical entity(medicinal substance), quality of the device needs to be assessed. The data requirements would depend on the kind of medicine. For example: Device that incorporates a prescription medicine must provide quality data in the format of Common Technical Document –Module 3. For example: DMF/PMF and /or EDQM Certificate of Suitability. The medicinal manufacturer may choose to provide either the DMF/PMF and /or EDQM Certificate of Suitability. Based on the option chosen, the requirements vary. Refer to guidance 11 of ARGPM for additional details(https://www.tga.gov.au/publication/australian-regulatory-guidelines-prescription-medicines-argpm ). For non-new chemical entity (eg. medicinal substances that are contained in other product/s registered/included in the ARTG) If the medicinal substance is manufactured by the same manufacturer (as other product/s already in the ARTG): • Confirmation of the substance quality • ARTG No. of the product/s containing the same medicinal substance • TGA DMF/PMF Reference No., if available (a letter from the medicinal substance manufacturer to allow TGA access to the DMF/PMF/EDQM Certificate of Suitability and confirmation that the quality of the medicinal substance contained in the device has not been changed since the last update of the DMF/PMF submitted to the TGA are required) Confirmation that the quality of the medicinal substance contained in the device is the same quality as the medicinal substance supplied for the product(s) registered/included in the ARTG If the medicinal substance is manufactured by a different manufacturer: • Quality data (eg. CTD Module 3, DMF/PMF and /or EDQM Certificate of Suitability with supporting data described in Guidance 11 of ARGPM)
CH3.5.02.1 | IMDRF | 3 | [Study description, study identifier, date of initiation] | CR | When any testing/study is conducted in support of the submission under the parent heading subject, this is required for each study/test. See information on Numbering of Headings in IMDRF ToC document
CH3.5.02.1.1 | IMDRF | 4 | Summary | CR | If applicable | Same as IMDRF content
CH3.5.02.1.2 | IMDRF | 4 | Full Report | CR | If applicable | Same as IMDRF content
CH3.5.02.1.3 | Regional (USFDA) | 4 | Statistical Data | NR
CH3.5.03 | IMDRF | 2 | Electrical Systems: Safety, Mechanical and Environmental Protection, and Electromagnetic Compatibility | CR | If applicable | In addition to IMDRF content,where applicable, testing certificate and accreditation status of laboratories used for testing in accordance with relevant applicable standards needs to be provided.
CH3.5.03.1 | IMDRF | 3 | [Study description, study identifier, date of initiation] | CR | When any testing/study is conducted in support of the submission under the parent heading subject, this is required for each study/test. See information on Numbering of Headings in IMDRF ToC document
CH3.5.03.1.1 | IMDRF | 4 | Summary | CR | If applicable | Same as IMDRF content
CH3.5.03.1.2 | IMDRF | 4 | Full Report | CR | If applicable | For all electronic and electric devices, Validation & Verification testing to demonstrate compliance with IEC 60601-1 and IEC 60601-1-2 or equivalent or better standard needs to be provided.
CH3.5.03.1.3 | Regional (USFDA) | 4 | Statistical Data | NR
CH3.5.04 | IMDRF | 2 | Radiation Safety | CR | If applicable | Where applicable, testing certificate and accreditation status of laboratories used for testing in accordance with relevant applicable standards needs to be provided.
CH3.5.04.1 | IMDRF | 3 | [Study description, study identifier, date of initiation] | CR | When any testing/study is conducted in support of the submission under the parent heading subject, this is required for each study/test. See information on Numbering of Headings in IMDRF ToC document
CH3.5.04.1.1 | IMDRF | 4 | Summary | CR | If applicable | Same as IMDRF content
CH3.5.04.1.2 | IMDRF | 4 | Full Report | CR | If applicable | Same as IMDRF content
CH3.5.04.1.3 | Regional (USFDA) | 4 | Statistical Data | NR
CH3.5.05 | IMDRF | 2 | Software/Firmware | CR | If applicable | Same as IMDRF content
CH3.5.05.01 | IMDRF | 3 | Software/Firmware Description | CR | If applicable | Same as IMDRF content
CH3.5.05.02 | IMDRF | 3 | Hazard Analysis | CR | If applicable | Same as IMDRF content
CH3.5.05.03 | IMDRF | 3 | Software Requirement Specification | CR | If applicable | Same as IMDRF content
CH3.5.05.04 | IMDRF(EU,HC,JP,USFDA,TGA) | 3 | Architecture Design Chart | CR | If applicable | Same as IMDRF content
CH3.5.05.05 | IMDRF(EU,HC,JP,USFDA,TGA) | 3 | Software Design Specification | CR | If applicable | Same as IMDRF content
CH3.5.05.06 | IMDRF | 3 | Traceability Analysis | CR | If applicable | Same as IMDRF content
CH3.5.05.07 | IMDRF | 3 | Software Life Cycle Process Description | CR | If applicable | Same as IMDRF content
CH3.5.05.08 | IMDRF | 3 | Software Verification and Validation | CR | If applicable | Where applicable, testing certificate and accreditation status of laboratories used for testing in accordance with relevant applicable standards needs to be provided.
CH3.5.05.08.1 | IMDRF | 4 | [Study description, study identifier, date of initiation] | CR | When any testing/study is conducted in support of the submission under the parent heading subject, this is required for each study/test. See information on Numbering of Headings in IMDRF ToC document
CH3.5.05.08.1.1 | IMDRF | 5 | Summary | CR | If applicable | Same as IMDRF content
CH3.5.05.08.1.2 | IMDRF | 5 | Full Report | CR | If applicable | For all devices incorporating software, evidence of compliance with IEC 62304 and IEC 62366 or equivalent or better standard is required.
CH3.5.05.08.1.3 | Regional (USFDA) | 5 | Statistical Data | NR
CH3.5.05.09 | IMDRF | 3 | Revision Level History | CR | If applicable | Same as IMDRF content
CH3.5.05.10 | IMDRF | 3 | Unresolved Anomalies (Bugs or Defects) | CR | If applicable | Same as IMDRF content
CH3.5.06 | IMDRF | 2 | Biocompatibility and Toxicology Evaluation | CR | If applicable | In addition to IMDRF content, in regards to toxicological evaluation; If the device incorporates a new chemical entity, a full toxicity profile needs to be assessed. The data submitted depends on the kind of medicine. For example: Device that incorporates a prescription medicine must provide toxicity data in the format of Common Technical Document – Module 4.
CH3.5.06.1 | IMDRF | 3 | [Study description, study identifier, date of initiation] | CR | When any testing/study is conducted in support of the submission under the parent heading subject, this is required for each study/test. See information on Numbering of Headings in IMDRF ToC document
CH3.5.06.1.1 | IMDRF | 4 | Summary | CR | If applicable | In regards to biocompatibility testing, a biological safety summary report (if available) in accordance with the current version of ISO 10993-1. If the current version is not used, justification of why use of the superseded version will not have detrimental impact on safety and effectiveness.
CH3.5.06.1.2 | IMDRF | 4 | Full Report | CR | If applicable | Full set of biological safety test reports (ISO 10993 or equivalent or better ) needs to be provided.
CH3.5.06.1.3 | Regional (USFDA) | 4 | Statistical Data | NR
CH3.5.07 | IMDRF | 2 | Non-Material-Mediated Pyrogenicity | CR | If applicable | Same as IMDRF content
CH3.5.07.1 | IMDRF | 3 | [Study description, study identifier, date of initiation] | CR | When any testing/study is conducted in support of the submission under the parent heading subject, this is required for each study/test. See information on Numbering of Headings in IMDRF ToC document
CH3.5.07.1.1 | IMDRF | 4 | Summary | CR | If applicable | Same as IMDRF content
CH3.5.07.1.2 | IMDRF | 4 | Full Report | CR | If applicable
CH3.5.07.1.3 | Regional (USFDA) | 4 | Statistical Data | NR
CH3.5.08 | IMDRF | 2 | Safety of Materials of Biological Origin (human/animal) | CR | If applicable | In addition to IMDRF content, details of the QMS records of the assessment and control of the suppliers of these materials. If device incorporates material of animal origin, the information in regards to source and manufacturing process of materials along with risk minimisation associated with transmission of pathogens is required. If device incorporates material of microbial/recombinant origin, the following information needs to be provided: Microbial origin • Microbial species – identification and cell bank qualification • Composition of fermentation and growth media by identifying all components along with their origin, suppliers, specifications and Certificates of Analysis. Recombinant origin • Identification and source of nucleotide sequence coding • Source of expression construct or host animals • Composition of fermentation of growth media including identification of all components along with their origin, suppliers, specifications and Certificates of Analysis.
CH3.5.08.1 | Regional (ANVISA,HC) | 3 | Certificates | NR
CH3.5.08.2 | IMDRF | 3 | [Study description, study identifier, date of initiation] | CR | When any testing/study is conducted in support of the submission under the parent heading subject, this is required for each study/test. See information on Numbering of Headings in IMDRF ToC document
CH3.5.08.2.1 | IMDRF | 4 | Summary | CR | If applicable | Same as IMDRF content
CH3.5.08.2.2 | IMDRF | 4 | Full Report | CR | If applicable | Same as IMDRF content
CH3.5.08.2.3 | Regional (USFDA) | 4 | Statistical Data | NR
CH3.5.09 | IMDRF | 2 | Sterilization Validation | -
CH3.5.09.1 | IMDRF | 3 | End-User Sterilization | CR | If applicable | Same as IMDRF content
CH3.5.09.1.1 | IMDRF | 4 | [Study description, study identifier, date of initiation] | CR | When any testing/study is conducted in support of the submission under the parent heading subject, this is required for each study/test. See information on Numbering of Headings in IMDRF ToC document
CH3.5.09.1.1.1 | IMDRF | 5 | Summary | CR | If applicable | Same as IMDRF content
CH3.5.09.1.1.2 | IMDRF | 5 | Full Report | CR | If applicable | Same as IMDRF content
CH3.5.09.1.1.3 | Regional (USFDA) | 5 | Statistical Data | NR
CH3.5.09.2 | IMDRF | 3 | Manufacturer Sterilization | CR | If applicable | Along with the IMDRF content for terminally sterilised products, information showing that the process has been physically and microbiologically validated to a SAL of 10-6. This should include: • The method of sterilisation used and the parameters (validation and routine) • A clear statement as to the standards being applied and/or details of any alternative sterilisation validation method used, if not specified in a recognised standard • The pre-sterilisation bioburden limit, the bioburden method used and validation of the method • A report describing the initial validation of the process and information/reports on revalidations carried out • Method of batch release ie BI release, sterility testing, dosimetric release or parametric release For products that are manufactured using sterile filtration and aseptic filling or aseptically manufactured: • bioburden information including pre-sterilisation bioburden limits and for extended processing times, evidence to show that microbiological quality (before sterilising filtration) and sterility (after sterilising filtration) is not compromised • parameters of sterilisation processes applied to the containers and closures and evidence to show that these processes have been physically and microbiologically validated to a SAL of 10-6 • details of filter integrity testing and information to show that the sterilising filter has been validated for bacterial retention in the presence of the product • statements of maximum permitted processing times during manufacture (holding, storage and filling times) • media fill studies to validate the aseptic manufacturing process. Media fill studies should be conducted under worst case conditions including maximum processing and filling times and should include simulation of all aseptic manufacturing processes • details of batch release sterility testing.
CH3.5.09.2.1 | IMDRF | 4 | [Study description, study identifier, date of initiation] | CR | When any testing/study is conducted in support of the submission under the parent heading subject, this is required for each study/test. See information on Numbering of Headings in IMDRF ToC document
CH3.5.09.2.1.1 | IMDRF | 5 | Summary | CR | If applicable | Same as IMDRF content
CH3.5.09.2.1.2 | IMDRF | 5 | Full Report | CR | If applicable | Full sterilisation validation reports of the sterilisation methods such as gamma, steam or ethylene oxide sterilisation methods conducted in accordance with the following - ISO 11137, ISO 17665, ISO 11135 or equivalent or better standards respectively needs to be provided.
CH3.5.09.2.1.3 | Regional (USFDA) | 5 | Statistical Data | NR
CH3.5.09.3 | IMDRF | 3 | Residual Toxicity | CR | If applicable
CH3.5.09.3.1 | IMDRF | 4 | [Study description, study identifier, date of initiation] | CR | When any testing/study is conducted in support of the submission under the parent heading subject, this is required for each study/test. See information on Numbering of Headings in IMDRF ToC document | Same as IMDRF content
CH3.5.09.3.1.1 | IMDRF | 5 | Summary | CR | If applicable | Same as IMDRF content
CH3.5.09.3.1.2 | IMDRF | 5 | Full Report | CR | If applicable | Same as IMDRF content
CH3.5.09.3.1.3 | Regional (USFDA) | 5 | Statistical Data | NR
CH3.5.09.4 | IMDRF | 3 | Cleaning and Disinfection Validation | CR | If applicable | Same as IMDRF content
CH3.5.09.4.1 | IMDRF | 3 | [Study description, study identifier, date of initiation] | CR | When any testing/study is conducted in support of the submission under the parent heading subject, this is required for each study/test. See information on Numbering of Headings in IMDRF ToC document
CH3.5.09.4.1.1 | IMDRF | 4 | Summary | CR | If applicable | Same as IMDRF content
CH3.5.09.4.1.2 | IMDRF | 4 | Full Report | CR | If applicable | Same as IMDRF content
CH3.5.09.4.1.3 | Regional (USFDA) | 5 | Statistical Data | NR
CH3.5.09.5 | IMDRF (ANVISA, USFDA) | 3 | Reprocessing of Single Use Devices, Validation Data | NR | If applicable
CH3.5.09.5.1 | IMDRF (ANVISA, USFDA) | 3 | [Study description, study identifier, date of initiation] | NR | When any testing/study is conducted in support of the submission under the parent heading subject, this is required for each study/test. See information on Numbering of Headings in IMDRF ToC document
CH3.5.09.5.1.1 | IMDRF (ANVISA, USFDA) | 4 | Summary | NR | If applicable
CH3.5.09.5.1.2 | IMDRF (ANVISA, USFDA) | 4 | Full Report | NR | If applicable
CH3.5.09.5.1.3 | Regional (USFDA) | 5 | Statistical Data | NR
CH3.5.10 | IMDRF | 2 | Animal Testing | CR | If applicable | Same as IMDRF content
CH3.5.10.1 | IMDRF | 3 | [Study description, study identifier, date of initiation] | CR | When any testing/study is conducted in support of the submission under the parent heading subject, this is required for each study/test. See information on Numbering of Headings in IMDRF ToC document
CH3.5.10.1.1 | IMDRF | 4 | Summary | CR | If applicable | Same as IMDRF content
CH3.5.10.1.2 | IMDRF | 4 | Full Report | CR | If applicable | Same as IMDRF content
CH3.5.10.1.3 | Regional (USFDA) | 4 | Statistical Data | NR
CH3.5.11 | IMDRF | 2 | Human Factors/Usability | CR | If applicable | Same as IMDRF content
CH3.5.11.1 | IMDRF | 3 | [Study description, study identifier, date of initiation] | CR | When any testing/study is conducted in support of the submission under the parent heading subject, this is required for each study/test. See information on Numbering of Headings in IMDRF ToC document
CH3.5.11.1.1 | IMDRF | 4 | Summary | CR | If applicable | Same as IMDRF content
CH3.5.11.1.2 | IMDRF | 4 | Full Report | CR | If applicable | Same as IMDRF content
CH3.5.11.1.3 | Regional (USFDA) | 4 | Statistical Data | NR
CH3.6 | IMDRF (HC, USFDA) | 1 | Non-clinical Bibliography | NR
CH3.7 | IMDRF | 1 | Expiration Period and Package Validation | R | For devices that do not have an expiration period, information regarding the estimated mean lifetime is required | In addition to IMDRF content, for devices that do not have an expiration period (e.g. electromedical equipment or other devices of multiple use), information regarding the estimated mean “lifetime”. This mean “lifetime” can be indicated as number of procedures to be performed with the device and/or its accessories, as a period of time or any other means of appropriate quantification.
CH3.7.1 | IMDRF | 2 | Product Stability | R |  | In addition to IMDRF content, if applicable, product stability shall also include: In use stability data, containing details and evidence supporting the stability during actual routine use of the device (real or simulated); and Shipping stability containing details and evidence supporting the tolerance of device components to the anticipated shipping conditions.
CH3.7.1.1 | IMDRF | 3 | [Study description, study identifier, date of initiation] | CR | When any testing/study is conducted in support of the submission under the parent heading subject, this is required for each study/test. See information on Numbering of Headings in IMDRF ToC document
CH3.7.1.1.1 | IMDRF | 4 | Summary | CR | If applicable | Same as IMDRF content
CH3.7.1.1.2 | IMDRF | 4 | Full Report | CR | If applicable | Same as IMDRF content
CH3.7.1.1.3 | Regional (USFDA) | 4 | Statistical Data | NR
CH3.7.2 | IMDRF | 2 | Package Validation | CR | If applicable | Same as IMDRF content
CH3.7.2.1 | IMDRF | 3 | [Study description, study identifier, date of initiation] | CR | When any testing/study is conducted in support of the submission under the parent heading subject, this is required for each study/test. See information on Numbering of Headings in IMDRF ToC document | Same as IMDRF content
CH3.7.2.1.1 | IMDRF | 4 | Summary | CR | If applicable | Same as IMDRF content
CH3.7.2.1.2 | IMDRF | 4 | Full Report | CR | If applicable | Same as IMDRF content
CH3.7.2.1.3 | Regional (USFDA) | 4 | Statistical Data | NR
CH3.8 | IMDRF | 1 | Other non-clinical Evidence | CR | When information is requested by the regulator (through guidance documents or other communication) but does not belong in any of the other non-clinical evidence headings. | For example, Testing to demonstrate compatibility of devices under MR conditions should be provided if MR compatability is claimed
CH3.8.1 | IMDRF | 2 | [Study description, study identifier, date of initiation] | CR | When any testing/study is conducted in support of the submission under the parent heading subject, this is required for each study/test. See information on Numbering of Headings in IMDRF ToC document
CH3.8.1.1 | IMDRF | 3 | Summary | CR | If applicable | Same as IMDRF content
CH3.8.1.2 | IMDRF | 3 | Full Report | CR | If applicable | Where applicable, for all passive MR conditional and safe devices Validation & Verification testing to demonstrate compliance with ASTM F2119, ASTM F2503, ASTM F2182, ASTM F2052, ASTM F2213 or equivalent or better standard. - For all active MR conditional devices Validation & Verification testing to demonstrate compliance with ISO 14708, ISO/TS 10974, ASTM F2119, ASTM F2503, ASTM F2182, ASTM F2052, ASTM F2213 or equivalent or better standard.
CH3.8.1.3 | Regional (USFDA) | 3 | Statistical Data | NR
CHAPTER 4 – CLINICAL EVIDENCE
CH4.1 | IMDRF | 1 | Chapter ToC | R |  | Same as IMDRF content
CH4.2 | IMDRF | 1 | Overall Clinical Evidence Summary | CR | Required unless this information is presented in the Clinical Evaluation Report
CH4.2.1 | IMDRF (TGA, EU, JP) | 2 | Clinical Evaluation Report | R |  | A clinical evaluation report (CER) written by an expert in the relevant field ensuring: the CER contains an objective critical evaluation of all of the clinical data submitted in relation to the device, with particular emphasis on whether and how the clinical data demonstrates safety and performance of the device for the intended purpose/use claimed. Please note that the manufacturer of the medical device must ensure that the clinical data is evaluated by a competent clinical expert.
CH4.2.2 | IMDRF | 2 | Device Specific Clinical Trials | -
CH4.2.2.1 | IMDRF | 3 | [Trial description, protocol #, date of initiation] | CR | When any testing/study is conducted in support of the submission under the parent heading subject, this is required for each study/test. See information on Numbering of Headings in IMDRF ToC document
CH4.2.2.1.1 | IMDRF | 4 | Clinical Trial Synopsis | CR | If applicable | Same as IMDRF content
CH4.2.2.1.2 | IMDRF | 4 | Clinical Trial Report | CR | If applicable | Where applicable, clinical trial data including all pivotal clinical study reports in full can be provided.
CH4.2.2.1.3 | Regional (USFDA) | 4 | Clinical Trial Data | NR
CH4.2.3 | IMDRF | 2 | Clinical Literature Review and Other Reasonable Known Information | CR | When a literature review is conducted in support of the submission. | In addition to IMDRF content, if Clinical literature review is method of evidence chosen, provide: • a description of the search strategy in sufficient detail such that the search can be replicated (e.g. document databases searched, search terms used, and inclusion and exclusion criteria applied); • pivotal publication(s) in full; and • sufficient detail for other publications to enable a technical assessment of the data, i.e. study design, methodology, statistics and outcome measures is required.
CH4.3 | Regional (USFDA) | 1 | IRB Approved Informed Consent Forms | NR
CH4.4 | Regional (USFDA) | 1 | Investigators Sites and IRB contact information | NR
CH4.5 | IMDRF | 1 | Other Clinical Evidence | CR | When information is requested by the regulator (through guidance documents or other communication) but does not belong in any of the other clinical evidence headings.
CH4.5.1 | IMDRF | 2 | [Study description, study identifier, date of initiation] | CR | When any testing/study is conducted in support of the submission under the parent heading subject, this is required for each study/test. See information on Numbering of Headings in IMDRF ToC document
CH4.5.1.1 | IMDRF | 3 | Summary | CR | If applicable | Same as IMDRF content
CH4.5.1.2 | IMDRF | 3 | Full Report | CR | If applicable | Same as IMDRF content
CH4.5.1.3 | Regional (USFDA) | 3 | Statistical Data | NR
CHAPTER 5 – LABELLING AND PROMOTIONAL MATERIAL
CH5.01 | IMDRF | 1 | Chapter ToC | R |  | Same as IMDRF content
CH5.02 | IMDRF, RF | 1 | Product/Package Labels | R |  | In addition to IMDRF content, the labels and instructions for use (including any package inserts) must a) meet the requirements of Essential Principle 13 b) be in English and legible when viewed on screen and printed c) include the Australian sponsor’s contact details to meet Regulation 10.2 If the applicant is including draft labels, artist impression or mock-up labels, the applicant needs to provide: a) the mock-up as full size suitable for A3 printing b) a statement as to where and how the batch/serial number/ date of manufacture/expiry date/ will be displayed.
CH5.03 | IMDRF, RF | 1 | Package Insert/Instructions for Use | R |  | In addition to IMDRF content, the labels and instructions for use (including any package inserts) must d) meet the requirements of Essential Principle 13 e) be in English and legible when viewed on screen and printed f) include the Australian sponsor’s contact details to meet Regulation 10.2 If the applicant is including draft labels, artist impression or mock-up labels, the applicant needs to provide: c) the mock-up as full size suitable for A3 printing d) a statement as to where and how the batch/serial number/ date of manufacture/expiry date/ will be displayed.
CH5.04 | IMDRF, RF (EU, Anvisa) | 1 | e-labelling | NR
CH5.05 | IMDRF (ANVISA, HC, TGA, USFDA) | 1 | Physician Labelling | CR | If applicable | Same as IMDRF content
CH5.06 | IMDRF (ANVISA, HC, USFDA) | 1 | Patient Labelling | NR | If applicable | Same as IMDRF content
CH5.07 | IMDRF (ANVISA, HC, TGA, USFDA) | 1 | Technical/Operator Manual | CR | If applicable | Same as IMDRF content
CH5.08 | Regional (ANVISA, HC, TGA) | 1 | Patient File Stickers/Cards and Implant Registration Cards | CR | When required or referenced as a risk control measure in the risk file | Same as IMDRF content
CH5.09 | Regional (HC,TGA) | 1 | Product Brochures | CR | Draft product brochures/advertising material available at the time of application | Same as IMDRF content
CH5.10 | IMDRF | 1 | Other Labelling and Promotional Material | CR | If applicable
CHAPTER 6A – QUALITY MANAGEMENT SYSTEM PROCEDURES
CH6A.1 | Regional (USFDA) | 1 | Cover Letter | NR | Cover letter required under Chapter 1
CH6A.2 | IMDRF (JP, TGA, USFDA) | 1 | Chapter ToC | R |  | Same as IMDRF content
CH6A.3 | IMDRF (JP, TGA, USFDA) | 1 | Administrative | -
CH6A.3.1 | IMDRF (JP, TGA, USFDA) | 2 | Product Descriptive Information | R | A list or brief description of all devices covered under the application; comprehensive product information is covered under Chapter 2 | Same as IMDRF content
CH6A.3.2 | IMDRF, RF (ANVISA, HC, JP, TGA, USFDA) | 2 | General Manufacturing Information | R |  | In addition to IMDRF content, it should also provide information in regards to the manufcaturing stages for each device.
CH6A.3.3 | IMDRF, RF (TGA, USFDA) | 2 | Required Forms | CR | When information is requested by the regulator (through guidance documents or other communication).
CH6A.4 | IMDRF, RF (TGA, USFDA) | 1 | Quality management system procedures | R |  | Same as IMDRF content
CH6A.5 | IMDRF, RF (TGA, USFDA) | 1 | Management responsibilities procedures | R |  | Same as IMDRF content
CH6A.6 | IMDRF, RF (TGA, USFDA) | 1 | Resource management procedures | R |  | Same as IMDRF content
CH6A.7 | IMDRF, RF (TGA, USFDA) | 1 | Product realization procedures | R |  | Same as IMDRF content
CH6A.7.1 | IMDRF, RF (TGA, USFDA) | 2 | Design and development procedures | CR | Required for applications under Part 1 - Full quality assurance procedures | Same as IMDRF content
CH6A.7.2 | IMDRF, RF (TGA, USFDA) | 2 | Purchasing procedures | R |  | Same as IMDRF content
CH6A.7.3 | IMDRF, RF (TGA, USFDA) | 2 | Production and service controls procedures | CR | Required for applications under Part 1 or Part 4 | Same as IMDRF content
CH6A.7.4 | IMDRF, RF (TGA, USFDA) | 2 | Control of monitoring and measuring devices procedures | CR | Required for applications under Part 1 or Part 4 | Same as IMDRF content
CH6A.8 | IMDRF, RF (TGA, USFDA) | 1 | QMS measurement, analysis and improvement procedures | R | Required for applications under Part 1 or Part 4 | In addition to IMDRF content, the following should also be included in this section: Procedures for a post-market monitoring system (Schedule 3, Part 1.4(3) or Part 4.4(3) of The Regulations). Procedures to notify the Secretary, in writing, of the proposed change; and arrange for assessment of the change by the Secretary to verify whether the system, as changed, meets the requirements. (Schedule 3, Clause 1.5(2) or 4.5(2) of The Regulations). Procedure for the issue and implementation of advisory notices and the notification of adverse events (Schedule 3, Part 1.4(3) or Part 4.4(3) of The Regulations and Uniform Recall Procedure for Therapeutic Goods).
CH6A.9 | IMDRF, RF (TGA, USFDA) | 1 | Other Quality System Procedures Information | CR | When information is requested by the regulator (through guidance documents or other communication) but does not belong in any of the other QMS Procedures headings.
CHAPTER 6B – QUALITY MANAGEMENT SYSTEM DEVICE SPECIFIC INFORMATION
CH6B.1 | IMDRF | 1 | Chapter ToC | R |  | Same as IMDRF content
CH6B.2 | IMDRF (TGA. USFDA) | 1 | Quality management system information | R |  | Same as IMDRF content
CH6B.3 | IMDRF (TGA. USFDA) | 1 | Management responsibilities information | CR | Any device specific procedures/records if applicable | Same as IMDRF content
CH6B.4 | IMDRF (TGA. USFDA) | 1 | Resource management information | CR | Any device specific procedures/records if applicable | Same as IMDRF content
CH6B.5 | Regional (HC) | 1 | Device Specific Quality Plan | NR
CH6B.6 | IMDRF (TGA. USFDA) | 1 | Product realization information | R |  | Same as IMDRF content
CH6B.6.1 | Regional (USFDA, Anvisa, TGA) | 2 | Design and development information | R | Required for devices undergoing design examination or type examination | Same as IMDRF content
CH6B.6.2 | IMDRF (TGA. USFDA) | 2 | Purchasing information | R |  | In addition to IMDRF content, List of suppliers of goods, services or outsourced processes that affect product conformity with requirements (critical suppliers); the goods, services or outsourced processes that they provide, and a description of how purchasing requirements are fulfilled for these suppliers
CH6B.6.3 | Regional (USFDA, HC, Anvisa, TGA) | 2 | Production and service controls information | CR | Any device specific procedures/records if applicable | Detailed Manufacturing Flow Diagram Summary of in-process acceptance activities for subject device Process Validation Master Plan List of processes that have not be validated For each process validation considered critical to the safety and effectiveness of the device: Protocols/Procedures for the validated process Process validation report The procedures for monitoring and controlling the process parameters of a validated process should be fully described. State the frequency of re-validation
CH6B.6.4 | IMDRF (TGA. USFDA) | 2 | Control of monitoring and measuring devices information | O |  | Same as IMDRF content
CH6B.7 | IMDRF (TGA. USFDA) | 1 | QMS measurement, analysis and improvement information | R | Any device specific procedures/records if applicable | In addition to IMDRF content, Adverse events (TGA /Other regulatory jurdisdictions reportable) Advisory notices (particularly recall, or field safety corrective action, notices) and alerts Trended complaints data for the previous three years from all markets to be provided. Any changes to the device should also be notified.
CH6B.8 | IMDRF (TGA. USFDA) | 1 | Other Device Specific Quality Management System Information | CR | When information is requested by the regulator (through guidance documents or other communication) but does not belong in any of the other QMS Device Specific Infomation headings. | Same as IMDRF content.